## Slide 1: Kinematics Presentation Tom
3_27_25

## Slide 2
~0.67V source, 16 bit enc, 8x ADC gain, 100 ohm div
Just Re-testing previous
3.3V source, 16 bit enc, 8x ADC gain, 50 ohm div

## Slide 3
3.3V source, 16 bit enc, 8x ADC gain, 50 ohm div
3.3V source from LVR and 2 x 10uF Caps, 16 bit enc, 8x ADC gain, 50 ohm div
3.3V source from LVR and 2 10uF Caps, 16 bit enc, 8x ADC gain, 40 ohm div (~15% of possible ADC range)

## Slide 4: Kinematics Presentation Tom
4_9_25

## Slide 5: Circuit Testing Summary
AD620 Instrumentation Amplifier Successfully configured to amplify the differential signal across the strain gauge and 2.048V Vref
LMC6840 Op-Amp Used to amplify the strain gauge voltage in a single-ended configuration. Provided stable output with low noise under 3.3V single-supply operation.
Power Supply Tried both 5V and 3.3V Low-Voltage Regulator (LVR)
Working Circuit Among several configurations tested, only one circuit demonstrated consistent and reliable performance. This setup will be the focus of further discussion and documentation.

## Slide 6
LVR +3.3V
GND
LMC6482, 75kohm/1kohm, total gain of 76
+26 mV to 52 mV
+0 mV
+1.98 V to 3.95V
LM040, 2.048V ref
-2.048 V
MCP3421 Differential Input PGA x2, 16 bit
(Alternative can use ADS111)
Arduino Uno I2C
ADC bits range from ~ -3000 to 20000 (as opposed to 5000 delta previously) of possible -32k to 32 k
GND
Strain Gauge (0.4 ohm – 0.8ohm)
50 ohm
x76
x2
Circuit V1 Option

## Slide 7
50 ohm
1k ohm
1 uF
0.1 uF
10 uF
LVR 3.3 V from 5V Arduino Uno
Actual Strain gauge about a foot away from this circuit

## Slide 8
LVR +3.3V
GND
+1.3 mV to 2.6 mV + Vref
~ - 1.95 mV (order precision resistor for this)+ Vref (average of input voltage range)
Microcontroller
GND
Strain Gauge (0.4 ohm – 0.8ohm)
1kohm (to limit current)
ADS112C04 or NAU7802 16-Bit, 4-Channel, 2-kSPS, Delta-Sigma ADC With I2C Interface, 128x Gain!
Circuit V2 Option (IDEAL)
X?
2.048V Vref
Using Differential ADC for Common Mode Rejection
NAU7802 Already does this but my output was very noisy!

## Slide 9: NAU7802 – use external 3.3V LVR and 2.048V ref (or slightly more) for AN- and 2.048V + strain gauge voltage for AN+
Serial.print("LDO voltage set to ");
switch (nau.getLDO()) {
case NAU7802_4V5: Serial.println("4.5V"); break;
case NAU7802_4V2: Serial.println("4.2V"); break;
case NAU7802_3V9: Serial.println("3.9V"); break;
case NAU7802_3V6: Serial.println("3.6V"); break;
case NAU7802_3V3: Serial.println("3.3V"); break;
case NAU7802_3V0: Serial.println("3.0V"); break;
case NAU7802_2V7: Serial.println("2.7V"); break;
case NAU7802_2V4: Serial.println("2.4V"); break;
case NAU7802_EXTERNAL: Serial.println("External"); break;

## Slide 10
3.3V source from LVR and 2 x 10uF Caps, 16 bit enc, 8x ADC gain, 50 ohm div
3.3V source from LVR and 2 x 10uF Caps, 16 bit enc, 2x ADC gain (using diff input), 76x gain from LMC6482 (across gauge), 50 ohm div with 0.4 ohm strain gauge (see circuit)
Bc of breadboarding, distance from strain gauge to circuit over a foot

## Slide 13
LVR +3.3V
GND
+1.3 mV to 2.6 mV + Vref
~ - 1.95 mV + Vref (average of input voltage range)
Microcontroller
GND
Strain Gauge (0.4 ohm – 0.8ohm)
1kohm (to limit current)
ADS112C04 16-Bit, 4-Channel, 2-kSPS, Delta-Sigma ADC With I2C Interface, 128x Gain!
Circuit V2 Option (IDEAL)
X?
2.048V Vref
Using Differential ADC for Common Mode Rejection

## Slide 14: Kinematics Presentation Tom
4_17_25

## Slide 15
Linear Fitting
Quadratic Fitting
Red Boxes = tests using full pos ADC range

## Slide 16
Linear Fitting
Quadratic Fitting

## Slide 17
Linear Fitting
Quadratic Fitting

## Slide 18
v
Velcro reapply
Velcro misalignment
0 , 10, 20 deg
Silicone Glue to Bender
Same Sample

## Slide 19
v
Velcro
Same Sample